## Slide 1: TIOLIDA - Toric IntraOcular Lens Implant Digital Alignment
Presenter: 陳孝寧
Email: r12922162@ntu.edu.tw
Tel: 0972064212

## Slide 2: Introduction

## Slide 3: Introduction | Cataract Surgery

## Slide 4: Introduction | Cataract Surgery
Surgery:
Femtosecond Laser-Assisted Cataract Surgery (FLACS): Higher operating cost, more limitations.

## Slide 5: Introduction | Cataract Surgery
Surgery:
By phacoemulsification (超音波乳化晶體摘除手術): Traditional method, lower cost and fewer limitations.

## Slide 6: Introduction | Cataract Surgery
Important factor:
During the operation, the size and position of the capsulorhexis (撕囊術) is important for the performance of IOL.
Small capsulorhexis can lead to anterior capsule fibrosis (水晶體前纖維化) and hyperopia shift (遠視偏移).
Large or asymmetric capsulorhexis can cause decentration, as well as subluxation (脫位) and posterior capsule opacification (水晶體後混濁) of IOL.
Accurate toric IOL alignment along the desired corneal meridian is a prerequisite for achieving optimal visual outcome. Approximately, one degree of misalignment reduces the astigmatism correction by 3.3%.

## Slide 7: Introduction | Cataract Surgery
Traditional manual marking:
https://www.youtube.com/watch?v=dAzb-wU7mgU
Marking Tool
Mark
Mark
Mark
Mark

## Slide 8: Introduction | Cataract Surgery
Drawbacks:
Color of mark may fade or disappear.
Contact marking may cause discomfort and risk of infection.
Accuracy of manual marking highly depends on surgeons.
Therefore computer-assisted IOL alignment is needed.

## Slide 9
Our Approach

## Slide 10: Experiment | Eye Center Detection

## Slide 11: Experiment | Eye Center Detection
Grayscale image
Binarized image
Input image
Read input image as grayscale image.
Binarize grayscale image by Otsu’s thresholding algorithm.

## Slide 12
Fit Ellipse
Calculate best fit ellipse from filtered contour.
Result:
    Works fine when the limbus boundary is complete and clear.
Issue:
    When part of the limbus is out of screen, or the limbus is not clear, the calculated ellipse will deform significantly.
NG case
OK case
Experiment | Eye Center Detection

## Slide 13
Primary trimming
Trim off the points on the edge of image from contour.
Result:
    Effectively remove irrelevant points on contour.
Without trimming
Trimmed and filtered contour
Experiment | Eye Center Detection

## Slide 14
Secondary trimming
Take 1st fit circle as reference, keep the points within a distance range and trim off the rest.
LSQ: Least Square
SVD: Singular Value Decomposition
1st trim circle
1st trim circle
1st trim contour
Experiment | Eye Center Detection

## Slide 15
Histogram equalization
Apply histogram equalization as preprocessing, to handle low-contrast input image.
Result:
    Successfully solved some low-contrast cases.
Without equalization
With equalization
Experiment | Eye Center Detection

## Slide 16
Truncate thresholding
Apply truncate thresholding as preprocessing, to handle overexposure input image.
Parameter: thres=135, pixel with intensity higher than threshold will be set to threshold.
Result:
    Successfully solved some overexposure cases.
Without truncated thresholding
With truncated thresholding
Experiment | Eye Center Detection

## Slide 17
Morphological operation - Opening
Apply opening operation on binarized image (with 11x11 pixels kernel), to reduce connection between limbus and tool, vessels, etc.
Result:
    Successfully increased limbus edge clarity.
Without opening
With opening
Experiment | Eye Center Detection

## Slide 18
Method: Hough circle transform
Select possible circle by voting and selecting local maximum in the Hough parameter space.
Concept illustration of Hough circle transform
Experiment | Eye Center Detection

## Slide 19
Method: Hough circle transform
Select possible circle by voting and selecting local maximum in the Hough parameter space.
Result:
    It seems not suitable for this application.
Hough circle transform result
Experiment | Eye Center Detection

## Slide 20
Model: U-Net++ with ResNet-101 backbone
Input: RGB images with size 3 channels x128 pixels x128 pixels
Classes:
0: Background
1: Limbus
Dataset:
Training Set: 5,362 video frames
Validation Set: 1,193 video frames
Number of epochs: 32
Loss function: Pixel-wise cross entropy loss.
Augmentation Strategies:
Random horizontal flip, Random vertical flip, Random shift-scale-rotation, Random brightness contrast, Random resized crop, Random Gaussian noise, Random color jitter.
Performance evaluation: IoU score of limbus area.
Experiment | Eye Center Detection
IoU: Intersection over Union

## Slide 21
Input size comparison
Experiment | Eye Center Detection

## Slide 22
Method 1: Use segmentation result directly to find contours.
Method 2: Use segmentation result as mask, then apply previous method.
Method 3: Use segmentation result as reference, apply distance thresholding on contour points.
Method 4: Use segmentation result as mask, apply previous method + RANSAC circle fit.
Experiment | Eye Center Detection
RANSAC: RANdom SAmple Consensus

## Slide 23
RANSAC circle fit:
1. Randomly choose 3 points from contour to form a circle, then count the points within threshold area.
2. Iterate several times and return the best circle.
Experiment | Eye Center Detection

## Slide 24
Previous result:
Experiment | Eye Center Detection
Hough circle

## Slide 25
Improved result:
Method 1
Method 3
Method 2
Method 4
Experiment | Eye Center Detection

## Slide 26
Evaluation:
Find contour and fitting ellipse from ground truth image.
Calculate cartesian distance between ellipse center of ground truth and detected circle center of each method.
Convert into mm by average limbus diameter (13.64 mm).
Experiment | Eye Center Detection

## Slide 27
Experiment | Eye Center Detection

## Slide 28
Experiment | Eye Center Detection

## Slide 29
Method: Apply LoFTR model to find matching pairs on ROI only instead of whole image.
Step: Extract polar area near limbus boundary.
ROI
Extracted ROI (Polar to Cartesian)
ROI: Region Of Interest
LoFTR: Local Feature TRansformer
Experiment | Rotation Tracking

## Slide 30
Method: Apply LoFTR model to find matching pairs on ROI only instead of whole image.
Step: Resize width to 1,080 pixels, find all matching pairs by LoFTR model.
This approach excludes irrelevant area and hence improves performance and accuracy.
ROI: Region Of Interest
LoFTR: Local Feature TRansformer
Experiment | Rotation Tracking

## Slide 31
Method: Apply LoFTR model to find matching pairs on ROI only instead of whole image.
Step: Estimate angle difference using RANSAC.
Set average confidence as threshold.
For each pair, if its confidence is above threshold, store its x-axis difference (pixels) into a list.
Apply RANSAC for 200 iterations (tentative), each iteration will pick 10% values from the list, calculate the mean value, then count the number of inlier values.
The mean value with maximum inlier value count is assigned to mean_pixel_diff.
Calculate angle difference by mean_pixel_diff/1,080, the error is expected to be less than 0.33°, 1,080 pixels: image width.
ROI: Region Of Interest
LoFTR: Local Feature TRansformer
RANSAC: RANdom SAmple Consensus
Experiment | Rotation Tracking

## Slide 32
Method: Apply LoFTR model to find matching pairs on ROI only instead of whole image.
Step: Draw the results and overlapped image.
Difficulty: the rotation of eyeball is not rigid, the deformation in each area is not constant, making it difficult to correctly evaluate the true accuracy of results.
ROI: Region Of Interest
LoFTR: Local Feature TRansformer
Experiment | Rotation Tracking

## Slide 33
Final result image
Experiment | Rotation Tracking

## Slide 34
Method: Toric intra-ocular lens implant digital alignment.
Steps: Detect limbus center -> Extract ROI around limbus edge -> Calculate rotational displacement -> Display alignment axis in current frame.
Results: Generate accurate alignment axis in surgery image to replace the locating dots in traditional method.
Summary

## Slide 35
Thanks for listening!

## Slide 36: Reference
Y. Zhai et al., "Computer-Aided Intraoperative Toric Intraocular Lens Positioning and Alignment During Cataract Surgery," in IEEE Journal of Biomedical and Health Informatics, vol. 25, no. 10, pp. 3921-3932, Oct. 2021.
S. Chaudhuri, S. Chatterjee, N. Katz, M. Nelson and M. Goldbaum, "Detection of blood vessels in retinal images using two-dimensional matched filters," in IEEE Transactions on Medical Imaging, vol. 8, no. 3, pp. 263-269, Sept. 1989, doi: 10.1109/42.34715.
https://ieee-dataport.org/open-access/cataracts
https://ieee-dataport.org/documents/cataract-surgery-dataset-eye-positioning-and-alignment#files
https://en.wikipedia.org/wiki/Otsu%27s_method
https://docs.opencv.org/4.x/d7/d4d/tutorial_py_thresholding.html
https://blog.csdn.net/hjxu2016/article/details/77833336
https://en.wikipedia.org/wiki/Roundness
https://shengyu7697.github.io/python-opencv-gray-to-rgb/

## Slide 37: Reference
https://docs.opencv.org/4.x/d7/d1b/group__imgproc__misc.html#ggaa9e58d2860d4afa658ef70a9b1115576a147222a96556ebc1d948b372bcd7ac59
https://en.wikipedia.org/wiki/Circle_Hough_Transform
https://docs.opencv.org/3.4/d4/d70/tutorial_hough_circle.html
https://pypi.org/project/circle-fit/